EZgeneTM  Blood Viral DNA/RNA Midiprep Kit

（BW-VR6512）

Contents

Kit Contents	2

Introduction	2

Storage and Stability	2

Before Starting	3

Important	3

Materials supplied by users	3

Protocol for Viral DNA/RNA extraction	3

Trouble Shooting Guide	4

Limited Use and warranty	5

- Kit Contents

| Catalog# | VR6512-00 | VR6512-01 | VR6512-02 |
| --- | --- | --- | --- |
| Preps | 1 | 10 | 25 |
| Buffer PLY | 22 mL | 210 mL | 500 mL |
| Wash Buffer * | 0.5 mL | 3 mL | 7mL |
| DEPC-Treated ddH2O | 110 µL | 1.1 mL | 2.6 mL |
| Mini Columns | 1 | 10 | 25 |
| Collection Tubes | 1 | 10 | 25 |
| Instruction Booklet | 1 | 1 | 1 |

*Add 2 mL (VR6512-00) or 12 mL (VR6512-01) or 28  mL (VR6512-02) 100% ethanol into each Wash Buffer bottle before use.

- Introduction

The EZgeneTM Viral DNA/RNA kit provides an easy and reliable method for isolating total viral DNA/RNA from plasma, serum, whole blood, urine and cell culture supernatant. This procedure has been tested for isolating nucleic acids from Hepatitis A, Hepatitis C and HIV. The isolated DNA/RNA can be used for PCR, RT-PCR and other downstream applications.

# Storage and Stability

All other components can be stored at room temperature(4-28℃). All kit components are guaranteed for 1 year from the date of production.

- Before Starting

Prepare all components and get all necessary materials ready by examining this instruction booklet and become familiar with each steps.

# Important

- Add 1% volume of β-mercaptoethanol to Buffer PLY before use and store at 4 oC.

- Add 2 mL (VR6512-00) or12 mL (VR6512-01) or 28 mL (VR6512-02) 100% ethanol into each Wash Buffer bottle before use.

# Materials supplied by users

- Tabletop microcentrifuge

- 1.5 mL sterile tubes

- 100% ethanol

Note: Perform all steps including centrifugation at room temperature

# Protocol for Viral DNA/RNA extraction

- Pipet 8-10 mL plasma, serum or other samples into a 50 mL tube and add 2 volumes of Buffer PLY.

Note: Ensure that β-mercaptoethanol has been added to Buffer PLY.

- Mix it thoroughly by vortexing and incubate for 15 min at room temperature.

- Add 0.5 volume 100% ethanol into the lysate (for example: 10 mL 100% ethanol for 20 mL lysate) and pipet 5 times to mix the solution.

- Insert a Mini column into a collection tube. Transfer the solution into Mini column and centrifuge at 13,000 rpm for 1 min. Discard the flow-through.

- Add 600 µL Wash Buffer to the column and centrifuge at 13,000 rpm for 1 min. Discard the flow-through.

- Repeat step 5.

- Discard the flow-through liquid ,reinsert the column to the collection tube. Centrifuge the empty column, with the lid open, at 13,000 rpm for 2 min.

Note: This step removes residual ethanol for optimal elution.

- Place the column to a RNase-free 1.5 mL tube, add 50-100 µL DEPC-Treated ddH2O to the column and centrifuge at 13,000 rpm for 1 min to elute the viral DNA/RNA (Store viral RNA at -20oC and viral DNA at 4 oC).

# Trouble Shooting Guide

| Problem | Possible reason | Suggested Improvement |
| --- | --- | --- |
| Low A260/A280 ratios | Protein contamination | Do a Phenol:Chloroform extraction. Loss of total RNA (up to 40%) should be expected. |
| Low A260/A280 ratios | Guanidine Thiocyanate contamination | Add 2.5 volumes of ethanol and 0.1M NaCl (final concentration) to precipitate RNA. Incubate for 30 min at -20°C. Centrifuge at 13,000 rpm for 15 min at 4°C. Resuspend the RNA pellet in DEPC-treated water. |
| Low Yield | RNA in sample degraded | Freeze samples immediately in liquid nitrogen and store at -70°C after collect it. |
| Low Yield | The binding capacity of the membrane in the spin column was exceeded | Use of too much sample exceeding the binding capacity of spin column will cause the decreasing of total RNA yield. |
| Low Yield | Ethanol not added to buffer | Add ethanol to the Wash Buffer. |
| Genomic DNA contamination | Too much total RNA sample was used in RT-PCR. | Reduce total RNA amount used in RT-PCR to 50-100 ng. |

# Limited Use and warranty

This product is intended for in vitro research use only. Not for use in human.This product is warranted to perform as described in its labeling and in BEIWO’s literature when used in accordance with instructions. No other warranties of any kind expressed or implied, including, without limitation, implied warranties of merchantability or fitness for a particular purpose, are provided by BEIWO. BEIWO’s sole obligation and purchaser’s exclusive remedy for breach of this warranty shall be, at the option of BEIWO, to replace the products, BEIWO shall have no liability for any direct, indirect, consequential, or incidental damage arising out of the use, the results of use, or the inability to use it product.

For technical support or learn more product information, please contact us or visit our website.

Contact Us: 400-115-2855

www.beiwobiomedical.com

Customer Support:

market@beiwobiomedical.com

Technical Support:

tech@beiwobiomedical.com